## Stack Flow Calcs
Rev 12/5/19, G. Allen |  |  |  | NOTE: | If Col. E Stack P is not in " Hg, the last term in H3 must be edited. | [5 empty] | These columns are static for a series of runs.
|  |  | Anemometer |  | Velocity |  | Velocity | Velocity
2020-04-14 | Stack T | Stack delta-p | Velocity | Stack P | m/s, cal. | Stack T | m/s | m/min | stack flow | stack flow | stack dia: | 6 | inches
Time (hh:mm) | deg F | "H20 absolute | m/sec raw | " Hg | corrected | deg C | at STP | STP | m3/min STP | CFM@STP | stack area: | 28.2744 | sq. inches |  |  |  |  | For Hontzsch ZS25/25 350GE/500/p10/ZG4 vane anemometer
8:50 | 75.2 |  | 1.46 | 29.675 | 1.4469 | 24 | 1.4399 | 86.3954 | 1.4539 | 51.344 | probe blockage area: | 2.63 | sq. inches | % block: | 9.3017 | Fc = | 8.3971 | This is the corrected blockage factor from EPA Nomograph 2, table VI. (October 1978) using a curve fit for actual blockage of 7.9 to 10.8% of area
9:00 | 201.6 |  | 1.73 | 29.675 | 1.739 | 94.2222 | 1.3996 | 83.9782 | 1.4132 | 49.9074 | [7 empty] | Raw blockage factor of 2.63 sq in. from O5 is based on sensor at the stack centerline and 1" circular probe diameter
9:10 | 296.5 |  | 1.75 | 29.675 | 1.7613 | 146.9444 | 1.2396 | 74.3771 | 1.2516 | 44.2016 | effective area: | 26.0841 | sq. in. -- using corrected blockage factor from U5
9:20 | 338 |  | 1.74 | 29.675 | 1.7501 | 170 | 1.1677 | 70.0593 | 1.179 | 41.6355 | effective area: | 0.016828 | m2
9:30 | 353.7 |  | 1.74 | 29.675 | 1.7501 | 178.7222 | 1.1451 | 68.7065 | 1.1562 | 40.8316
9:40 | 354.6 |  | 1.8 | 29.675 | 1.8175 | 179.2222 | 1.1879 | 71.2712 | 1.1994 | 42.3558 |  |  | Anemometer model number: | ZS25/25-350GE/500/P10/ZG4
9:50 | 358.5 |  | 1.81 | 29.675 | 1.8288 | 181.3889 | 1.1895 | 71.3726 | 1.2011 | 42.416 |  |  | Anemometer serial number: | mn20-13393
10:00 | 357.5 |  | 1.81 | 29.675 | 1.8288 | 180.8333 | 1.191 | 71.4599 | 1.2026 | 42.468 | Anemometer cal equation factors:
10:10 | 361.8 |  | 1.82 | 29.675 | 1.8401 | 183.2222 | 1.1921 | 71.5261 | 1.2037 | 42.5073 | B0 | B1 | B2 | Cal Date: | 2019-11-14
10:20 | 361.5 |  | 1.81 | 29.675 | 1.8288 | 183.0556 | 1.1852 | 71.1117 | 1.1967 | 42.261 | 0.16099 | 0.71124 | 0.11613
10:30 | 362.9 |  | 1.8 | 29.675 | 1.8175 | 183.8333 | 1.1759 | 70.5518 | 1.1873 | 41.9283 | actual = 0.161 + (0.7112 * observed) + (0.11613 * observed^2)
10:40 | 365.1 |  | 1.8 | 29.675 | 1.8175 | 185.0556 | 1.1727 | 70.3636 | 1.1841 | 41.8164
10:50 | 347.2 |  | 1.81 | 29.675 | 1.8288 | 175.1111 | 1.2062 | 72.3725 | 1.2179 | 43.0103 |  | Room BP at start of test: | 29.7 | " Hg
11:00 | 287.8 |  | 1.83 | 29.675 | 1.8515 | 142.1111 | 1.3182 | 79.0947 | 1.331 | 47.0052 |  | Room BP at end of test: | 29.65 | " Hg
11:10 | 285 |  | 1.85 | 29.675 | 1.8742 | 140.5556 | 1.3395 | 80.3687 | 1.3525 | 47.7623 |  | Avg BP: | 29.675 | " Hg
11:20 | 289.4 |  | 1.84 | 29.675 | 1.8628 | 143 | 1.3235 | 79.4106 | 1.3364 | 47.1929
11:30 | 282.6 |  | 1.85 | 29.675 | 1.8742 | 139.2222 | 1.3438 | 80.6286 | 1.3569 | 47.9168 | 1 " H20 = | 0.07348 | " Hg
11:40 | 285.7 |  | 1.85 | 29.675 | 1.8742 | 140.9444 | 1.3382 | 80.2932 | 1.3512 | 47.7175 | 1 mb = | 0.02953 | " Hg
11:50 | 286.5 |  | 1.85 | 29.675 | 1.8742 | 141.3889 | 1.3368 | 80.2071 | 1.3498 | 47.6663
12:00 | 288.2 |  | 1.84 | 29.675 | 1.8628 | 142.3333 | 1.3256 | 79.538 | 1.3385 | 47.2687
12:10 | 284.8 |  | 1.85 | 29.675 | 1.8742 | 140.4444 | 1.3398 | 80.3903 | 1.3528 | 47.7752
12:20 | 286.7 |  | 1.85 | 29.675 | 1.8742 | 141.5 | 1.3364 | 80.1856 | 1.3494 | 47.6535
12:30 | 289.8 |  | 1.8 | 29.675 | 1.8175 | 143.2222 | 1.2906 | 77.4356 | 1.3031 | 46.0193
12:40 | 288.2 |  | 1.85 | 29.675 | 1.8742 | 142.3333 | 1.3337 | 80.0247 | 1.3467 | 47.5579
12:50 | 291.4 |  | 1.85 | 29.675 | 1.8742 | 144.1111 | 1.3281 | 79.6836 | 1.3409 | 47.3552
13:00 | 259.8 |  | 1.85 | 29.675 | 1.8742 | 126.5556 | 1.3864 | 83.1847 | 1.3999 | 49.4359
13:10 | 237.4 |  | 1.83 | 29.675 | 1.8515 | 114.1111 | 1.4136 | 84.8157 | 1.4273 | 50.4051
13:20 | 237.7 |  | 1.84 | 29.675 | 1.8628 | 114.2778 | 1.4217 | 85.3 | 1.4355 | 50.693
13:30 | 239.8 |  | 1.83 | 29.675 | 1.8515 | 115.4444 | 1.4087 | 84.5245 | 1.4224 | 50.2321
13:40 | 236.7 |  | 1.84 | 29.675 | 1.8628 | 113.7222 | 1.4237 | 85.4225 | 1.4375 | 50.7658
13:50 | 238 |  | 1.83 | 29.675 | 1.8515 | 114.4444 | 1.4124 | 84.7427 | 1.4261 | 50.3618
14:00 | 239.3 |  | 1.84 | 29.675 | 1.8628 | 115.1667 | 1.4184 | 85.1047 | 1.4322 | 50.5769
14:10 | 242.6 |  | 1.83 | 29.675 | 1.8515 | 117 | 1.4031 | 84.1874 | 1.4167 | 50.0318
14:20 | 237.8 |  | 1.82 | 29.675 | 1.8401 | 114.3333 | 1.4041 | 84.2473 | 1.4177 | 50.0674
14:30 | 240.8 |  | 1.82 | 29.675 | 1.8401 | 116 | 1.3981 | 83.8863 | 1.4117 | 49.8528
14:40 | 239.4 |  | 1.83 | 29.675 | 1.8515 | 115.2222 | 1.4095 | 84.5729 | 1.4232 | 50.2609
14:50 | 236.9 |  | 1.82 | 29.675 | 1.8401 | 113.8333 | 1.4059 | 84.3562 | 1.4196 | 50.1321
15:00 | 238.5 |  | 1.83 | 29.675 | 1.8515 | 114.7222 | 1.4114 | 84.682 | 1.4251 | 50.3257
15:10 | 242 |  | 1.82 | 29.675 | 1.8401 | 116.6667 | 1.3957 | 83.7428 | 1.4093 | 49.7676
15:20 | 239.7 |  | 1.82 | 29.675 | 1.8401 | 115.3889 | 1.4003 | 84.0183 | 1.4139 | 49.9313
15:30 | 239.2 |  | 1.81 | 29.675 | 1.8288 | 115.1111 | 1.3927 | 83.5609 | 1.4062 | 49.6594
15:40 | 239.6 |  | 1.82 | 29.675 | 1.8401 | 115.3333 | 1.4005 | 84.0304 | 1.4141 | 49.9384
Note on O5: this is from physical dims of the probe

## Sheet1
ISS 2
| SCFH | SCFM
1 | 4.99 | 0.083167
2 | 19.98 | 0.333
3 | 34.99 | 0.58317
4 | 50.01 | 0.8335
5 | 65.02 | 1.0837
6 | 79.94 | 1.3323
7 | 95.05 | 1.5842
8 | 110.01 | 1.8335

## Instructions
For Hontzsch ZS25/25 350GE/500/p10/ZG4 vane anemometer
Enter or verify all values in cells with this highlight:
Blockage correction factor is constant for a given probe insertion distance. See anemometer SOP for more information.
Stack P is based on room BP minus the measured delta-P in the stack.
C3 is absolute Delta-P -- always a postive #.
"Stack velocity, m/s raw" is the flow read from the anemometer software.
If the Hontzsch anemometer is not in the centerline of a 6" stack, the blockage factor in O5 must be changed. | 11/14/19 TransCat calibration results:
| actual | observed | (m/s) | corrected reading
| 0.44 | 0.37 |  | 0.44004
| 0.59 | 0.56 |  | 0.5957
Revision history: | 0.79 | 0.77 |  | 0.7775
11/13/19 First release | 1 | 1.02 |  | 1.0073
11/14/19 improved documentation; no data changes | 1.99 | 1.95 |  | 1.9895
11/14/19 Fixed bad formula in K5 [Kelli] | 5 | 4.9 |  | 6.4343
11/14/19 Entered actual cal factors from 11/14/19 calibration and corrected the example info for that equation.
11/14/19 Fixed bad formula in F5 [Nick] | Without 5 lpm point:
11/15/19 Put check for 'out of curve range' in U5 | y = 0.160985 + 0.711244*x + 0.116128*x^2
11/29/19 Added more sig figures to cal equation | r2 = 1.000