## Stack Flow Calcs
Rev 12/5/19, G. Allen |  |  |  | NOTE: | If Col. E Stack P is not in " Hg, the last term in H3 must be edited. | [5 empty] | These columns are static for a series of runs.
|  |  | Anemometer |  | Velocity |  | Velocity | Velocity
2020-01-21 | Stack T | Stack delta-p | Velocity | Stack P | m/s, cal. | Stack T | m/s | m/min | stack flow | stack flow | stack dia: | 6 | inches
Time (hh:mm) | deg F | "H20 absolute | m/sec raw | " Hg | corrected | deg C | at STP | STP | m3/min STP | CFM@STP | stack area: | 28.2744 | sq. inches |  |  |  |  | For Hontzsch ZS25/25 350GE/500/p10/ZG4 vane anemometer
9:25 | 60.7 |  | 0 | 30.05 | 0.16099 | 15.9444 | 0.16675 | 10.0051 | 0.16837 | 5.9459 | probe blockage area: | 2.63 | sq. inches | % block: | 9.3017 | Fc = | 8.3971 | This is the corrected blockage factor from EPA Nomograph 2, table VI. (October 1978) using a curve fit for actual blockage of 7.9 to 10.8% of area
9:35 | 61.8 |  | 1.32 | 30.05 | 1.3022 | 16.5556 | 1.346 | 80.758 | 1.359 | 47.9937 | [7 empty] | Raw blockage factor of 2.63 sq in. from O5 is based on sensor at the stack centerline and 1" circular probe diameter
9:45 | 150 |  | 1.26 | 30.05 | 1.2415 | 65.5556 | 1.0975 | 65.8527 | 1.1082 | 39.1356 | effective area: | 26.0841 | sq. in. -- using corrected blockage factor from U5
9:55 | 230.6 |  | 1.18 | 30.05 | 1.1619 | 110.3333 | 0.90721 | 54.4329 | 0.91602 | 32.3489 | effective area: | 0.016828 | m2
10:05 | 260.2 |  | 1.1 | 30.05 | 1.0839 | 126.7778 | 0.81144 | 48.6865 | 0.81932 | 28.9339
10:15 | 270.8 |  | 1.09 | 30.05 | 1.0742 | 132.6667 | 0.79254 | 47.5523 | 0.80023 | 28.2599 |  |  | Anemometer model number: | ZS25/25-350GE/500/P10/ZG4
10:25 | 274.3 |  | 1.1 | 30.05 | 1.0839 | 134.6111 | 0.79585 | 47.7508 | 0.80357 | 28.3779 |  |  | Anemometer serial number: | mn20-13393
10:35 | 276.9 |  | 1.08 | 30.05 | 1.0646 | 136.0556 | 0.77892 | 46.7355 | 0.78648 | 27.7744 | Anemometer cal equation factors:
10:45 | 277.6 |  | 1.1 | 30.05 | 1.0839 | 136.4444 | 0.79228 | 47.537 | 0.79997 | 28.2508 | B0 | B1 | B2 | Cal Date: | 2019-11-14
10:55 | 278.1 |  | 1.1 | 30.05 | 1.0839 | 136.7222 | 0.79175 | 47.5048 | 0.79943 | 28.2316 | 0.16099 | 0.71124 | 0.11613
11:05 | 278.5 |  | 1.09 | 30.05 | 1.0742 | 136.9444 | 0.78427 | 47.0561 | 0.79188 | 27.965 | actual = 0.161 + (0.7112 * observed) + (0.11613 * observed^2)
11:15 | 279.3 |  | 1.07 | 30.05 | 1.055 | 137.3889 | 0.76939 | 46.1631 | 0.77685 | 27.4343
11:25 | 282.7 |  | 1.06 | 30.05 | 1.0454 | 139.2778 | 0.7589 | 45.5341 | 0.76627 | 27.0605 |  | Room BP at start of test: | 30.05 | " Hg
11:35 | 258.9 |  | 1.08 | 30.05 | 1.0646 | 126.0556 | 0.79844 | 47.9066 | 0.80619 | 28.4704 |  | Room BP at end of test: | 30.05 | " Hg
11:45 | 229.2 |  | 1.15 | 30.05 | 1.1325 | 109.5556 | 0.88601 | 53.1609 | 0.89461 | 31.593 |  | Avg BP: | 30.05 | " Hg
11:55 | 230.1 |  | 1.12 | 30.05 | 1.1032 | 110.0556 | 0.86201 | 51.7205 | 0.87037 | 30.737
12:05 | 226.1 |  | 1.13 | 30.05 | 1.113 | 107.8333 | 0.87468 | 52.4808 | 0.88317 | 31.1889 | 1 " H20 = | 0.07348 | " Hg
12:15 | 236.1 |  | 1.15 | 30.05 | 1.1325 | 113.3889 | 0.87722 | 52.6335 | 0.88574 | 31.2796 | 1 mb = | 0.02953 | " Hg
12:25 | 229.7 |  | 1.14 | 30.05 | 1.1227 | 109.8333 | 0.87773 | 52.6639 | 0.88625 | 31.2977
12:35 | 235.3 |  | 1.17 | 30.05 | 1.1521 | 112.9444 | 0.89344 | 53.6067 | 0.90212 | 31.8579
12:45 | 234.6 |  | 1.15 | 30.05 | 1.1325 | 112.5556 | 0.87912 | 52.7472 | 0.88765 | 31.3472
12:55 | 232.4 |  | 1.14 | 30.05 | 1.1227 | 111.3333 | 0.87431 | 52.4584 | 0.88279 | 31.1755
13:05 | 238.4 |  | 1.17 | 30.05 | 1.1521 | 114.6667 | 0.88948 | 53.3685 | 0.89811 | 31.7164
13:15 | 233.8 |  | 1.13 | 30.05 | 1.113 | 112.1111 | 0.86496 | 51.8979 | 0.87336 | 30.8424
13:25 | 233 |  | 1.14 | 30.05 | 1.1227 | 111.6667 | 0.87355 | 52.4129 | 0.88203 | 31.1485
13:35 | 229.1 |  | 1.13 | 30.05 | 1.113 | 109.5 | 0.87087 | 52.2522 | 0.87932 | 31.053
13:45 | 200.6 |  | 1.16 | 30.05 | 1.1423 | 93.6667 | 0.9324 | 55.9442 | 0.94145 | 33.2471
13:55 | 193.8 |  | 1.15 | 30.05 | 1.1325 | 89.8889 | 0.93403 | 56.0419 | 0.9431 | 33.3052
14:05 | 202.8 |  | 1.13 | 30.05 | 1.113 | 94.8889 | 0.90546 | 54.3274 | 0.91424 | 32.2863
14:15 | 196 |  | 1.12 | 30.05 | 1.1032 | 91.1111 | 0.90686 | 54.4114 | 0.91566 | 32.3362
14:25 | 197.1 |  | 1.13 | 30.05 | 1.113 | 91.7222 | 0.91332 | 54.7991 | 0.92218 | 32.5666
14:35 | 198 |  | 1.13 | 30.05 | 1.113 | 92.2222 | 0.91207 | 54.7241 | 0.92092 | 32.522
14:45 | 197 |  | 1.12 | 30.05 | 1.1032 | 91.6667 | 0.90548 | 54.3285 | 0.91426 | 32.2869
14:55 | 195.9 |  | 1.12 | 30.05 | 1.1032 | 91.0556 | 0.907 | 54.4197 | 0.9158 | 32.3411
15:05 | 192.7 |  | 1.13 | 30.05 | 1.113 | 89.2778 | 0.91948 | 55.1689 | 0.9284 | 32.7863
15:15 | 196.1 |  | 1.11 | 30.05 | 1.0935 | 91.1667 | 0.89875 | 53.9247 | 0.90747 | 32.0469
15:25 | 194.7 |  | 1.09 | 30.05 | 1.0742 | 90.3889 | 0.88474 | 53.0847 | 0.89333 | 31.5477
15:35 | 195.4 |  | 1.14 | 30.05 | 1.1227 | 90.7778 | 0.92371 | 55.4226 | 0.93267 | 32.9371
15:45 | 195.5 |  | 1.1 | 30.05 | 1.0839 | 90.8333 | 0.89161 | 53.4964 | 0.90026 | 31.7924
15:55 | 198 |  | 1.1 | 30.05 | 1.0839 | 92.2222 | 0.88822 | 53.293 | 0.89684 | 31.6715
16:05 | 191.7 |  | 1.13 | 30.05 | 1.113 | 88.7222 | 0.92089 | 55.2536 | 0.92983 | 32.8367
16:15 | 193 |  | 1.14 | 30.05 | 1.1227 | 89.4444 | 0.92711 | 55.6265 | 0.93611 | 33.0583
16:25 | 191.33 |  | 1.14 | 30.05 | 1.1227 | 88.5167 | 0.92949 | 55.7693 | 0.93851 | 33.1431
Note on O5: this is from physical dims of the probe

## Sheet1
ISS 2
| SCFH | SCFM
1 | 4.99 | 0.083167
2 | 19.98 | 0.333
3 | 34.99 | 0.58317
4 | 50.01 | 0.8335
5 | 65.02 | 1.0837
6 | 79.94 | 1.3323
7 | 95.05 | 1.5842
8 | 110.01 | 1.8335

## Instructions
For Hontzsch ZS25/25 350GE/500/p10/ZG4 vane anemometer
Enter or verify all values in cells with this highlight:
Blockage correction factor is constant for a given probe insertion distance. See anemometer SOP for more information.
Stack P is based on room BP minus the measured delta-P in the stack.
C3 is absolute Delta-P -- always a postive #.
"Stack velocity, m/s raw" is the flow read from the anemometer software.
If the Hontzsch anemometer is not in the centerline of a 6" stack, the blockage factor in O5 must be changed. | 11/14/19 TransCat calibration results:
| actual | observed | (m/s) | corrected reading
| 0.44 | 0.37 |  | 0.44004
| 0.59 | 0.56 |  | 0.5957
Revision history: | 0.79 | 0.77 |  | 0.7775
11/13/19 First release | 1 | 1.02 |  | 1.0073
11/14/19 improved documentation; no data changes | 1.99 | 1.95 |  | 1.9895
11/14/19 Fixed bad formula in K5 [Kelli] | 5 | 4.9 |  | 6.4343
11/14/19 Entered actual cal factors from 11/14/19 calibration and corrected the example info for that equation.
11/14/19 Fixed bad formula in F5 [Nick] | Without 5 lpm point:
11/15/19 Put check for 'out of curve range' in U5 | y = 0.160985 + 0.711244*x + 0.116128*x^2
11/29/19 Added more sig figures to cal equation | r2 = 1.000