Critical Criteria – Ozone

| Requirement (O3) | Frequency | Acceptable Range | Review Comments |
| --- | --- | --- | --- |
| Monitor | NA | Meets requirements listed in FRM/FEM designation |  |
| One-Point QC Check Single analyzer | Every 14 days | < ± 7.1% (percent difference) or < ± 1.5 ppb difference whichever is greater |  |
| Zero/Span Check | Every 14 days | Zero drift (24-hr) < ± 3.1 ppb Zero drift (>24hr-14 day) < ± 5.1 ppb Span drift < ± 7.1% |  |

Operational Criteria – Ozone

| Requirement (O3) | Frequency | Acceptable Range | Review Comments |
| --- | --- | --- | --- |
| Shelter Temperature Range | Daily (hourly values) | 20 to 30° C (hourly avg.) or per manufacturers specification if designated to a wider temp. range |  |
| Shelter Temperature Control | Daily (hourly values) | < 2.1° C SD over 24 hours |  |
| Shelter Temperature Device Check | Every 182 days and 2/calendar year | < ± 2.1° C of standard |  |
| Annual Performance Evaluation Single Analyzer (SLAMS/NCORE) | Every site every 365 days and 1/calendar year within period of monitor operation | Percent difference of audit levels 3-10 < ±15.1%, Audit levels 1 & 2 < ± 1.5 ppb difference or < ±15.1% |  |
| Annual Performance Evaluation Single Analyzer (PSD) | Quarterly | Percent difference of audit levels 3-10 < ±15.1%, Audit levels 1 & 2 < ± 1.5 ppb difference or < ±15.1% |  |
| Federal Audit (NPAP) (SLAMS/NCORE) | 20% of sites audited in calendar year | Audit levels 1 & 2 < ± 1.5 ppb difference, all other levels percent difference < ±10.1% |  |
| Verification / Calibration | Upon receipt/adjustment/repair/installation/moving and/or recalibration of higher level transfer standard Every 182 days and 2/calendar year if manual biweekly zero/span checks, or Every 365 days and 1/calendar year if continuous daily zero/span checks | All points < ± 2.1% or ≤ ±1.5 ppb difference of best-fit straight line, whichever is greater, and Slope 1 ± .05 |  |
| Zero Air / Zero Air Check | Every 365 days and 1/calendar year | Concentration below LDL |  |

Ozone Level 2 Standard

| Certification/recertification to Standard Reference Photometer (Level 1) | Every 365 days and 1/calendar year | Single point difference < ± 3.1% |  |
| --- | --- | --- | --- |
| Level 2 and greater transfer standard precision | Every 365 days and 1/calendar year | Standard Deviation less than 0.005 ppm or 3.0% whichever is greater |  |
| (if recertified via a transfer standard) | Every 365 days and 1/calendar year | Regression slope = 1.00 ± 0.03 and two intercepts are 0 ±3 ppb |  |

Ozone Transfer standard (Level 3 and greater)

| Qualification | Upon receipt of transfer standard | < ± 4.1% or < ±4 ppb (whichever is greater) |  |
| --- | --- | --- | --- |
| Certification | After qualification and upon receipt/ adjustment/repair | RSD of six slopes ≤ 3.7% Std. Dev. of 6 intercepts ≤ 1.5 |  |
| Recertification to higher level standard | Beginning and end of O3 season or every 182 days and 2/calendar year whichever less | New slope = ± 0.05 of previous and RSD of six slopes ≤ 3.7% Std. Dev. of 6 intercepts 1.5 |  |

Detection (FEM/FRMs)

| Noise | Every 365 days and 1/calendar year | ≤ 0.0025 ppm (standard range) ≤ 0.001 ppm (lower range) |  |
| --- | --- | --- | --- |
| Lower detectable level | Every 365 days and 1/calendar year | ≤ 0.005 ppm (standard range) ≤ 0.002 ppm (lower range) |  |

SYSTEMATIC CRITERIA – O3

| Requirement (O3) | Frequency | Acceptable Range | Review Comments |
| --- | --- | --- | --- |
| Standard Reporting Units | All data | ppm (final units in AQS) |  |
| Rounding convention for data reported to AQS | All data | 3 places after decimal with digits to right truncated |  |
| Completeness (SLAMS/NCORE) | 3-year Comparison | ≥ 90% (avg.) daily max available in ozone season with min of 75% in any one year |  |
| Completeness (SLAMS/NCORE) | 8-hour Average | ≥ if at least 6 of the hourly concentrations for the 8-hour period are available |  |
| Completeness (SLAMS/NCORE) | Valid Daily Max | ≥ if valid 8-hour averages are available for at least 13 of the 17 consecutive 8-hour periods starting from 7:00 a.m. to 11:00 p.m. |  |
| Completeness (PSD) | 1-year | ≥ 80% of each year |  |
| Completeness (PSD) | 8-hour Average | ≥ if at least 6 of the hourly concentrations for the 8-hour period are available |  |
| Completeness (PSD) | Valid Daily Max | ≥ 75% of the 24, 8-hour averages (18 of 24, 8-hour averages) |  |
| Sample Residence Time Verification | Every 365 days and 1/calendar year | ≤ 20 seconds |  |
| Sample Probe, Inlet, Sampling train | All sites | Borosilicate glass,(e.g. Pyrex®) or Teflon® (FEP/TFE) |  |
| Data Acquisition Systems | Digital or analog recording devices | Collection of continuous data (minimum of 1-minute values) |  |
| System Clock Verification | 1 / month | ≤ ± 1 minute |  |
| Siting | Every 365 days and 1/calendar year | Meeting siting criteria or waiver documented, (PSD per approved QAPP) |  |
| EPA Standard Ozone Reference Photometer (SRP) Recertification (Level 1) | Every 365 days and 1/calendar year | Regression sloe = 1.00 ± 0.01 and intercept < 0.003 ppm |  |
| Precision (using 1-point QC Checks) (SLAMS/NCore) | Calculated annually and as appropriate for design value estimates | 90% CL CV < 7.1% |  |
| Precision (using 1-point QC Checks) (PSD) | Calculated Quarterly | 90% CL CV < 7.1% |  |
| Bias (using 1-point QC Checks) (SLAMS/NCore) | Calculated annually and as appropriate for design value estimates | 95% CL < ± 7.1% |  |
| Bias (using 1-point QC Checks) (PSD) | Calculated Quarterly | 95% CL < ± 7.1% |  |
| Technical Systems Audits (SLAMS/NCore) | 1 / 3 years | Confirmation of adherence to FRM/FEM, SOPs, and QAPP or documented waiver |  |
| Technical Systems Audits (PSD) | Annually (within 1 month of startup and annually thereafter) | Confirmation of adherence to FRM/FEM, SOPs, and QAPP or documented waiver |  |
| Annual PE Primary QA Organization (PQAO) Evaluation | 1 / year | 95% of audit percent difference fall within the 1-point QC check 95% probability intervals at PQAO level of aggregation |  |