Critical Criteria – NO2

| Requirement (NO2) | Frequency | Acceptable Range | Review Comments |
| --- | --- | --- | --- |
| Sampler/Monitor | NA | Meets requirements listed in FRM/FEM designation |  |
| One-Point QC Check Single analyzer | Every 14 days | < ± 15.1% (percent difference) or < ± 1.5 ppb difference whichever is greater |  |
| Zero/Span Check | Every 14 days | Zero drift < ± 3.1 ppb (24-hr), < ± 5.1 ppb (>24hr-14 days) Span drift < ± 10.1% |  |
| Converter Efficiency | During multi-point calibrations, spans, and audits Every 14 days | 96% - 104.1% |  |

Operational Criteria – NO2

| Requirement (NO2) | Frequency | Acceptable Range | Review Comments |
| --- | --- | --- | --- |
| Shelter Temperature Range | Daily (hourly values) | 20 to 30° C (hourly avg.) or per manufacturers specification if designated to a wider temp. range |  |
| Shelter Temperature Control | Daily (hourly values) | < 2.1° C SD over 24 hours |  |
| Shelter Temperature Device Check | Every 182 days and 2/calendar year | < ± 2.1° C of standard |  |
| Annual Performance Evaluation Single Analyzer (SLAMS/NCORE) | Every site every 365 days and 1/calendar year | Percent difference of audit levels 3-10 < ± 15.1%, Audit levels 1&2 < ± 1.5 ppb difference or < ±15.1% |  |
| Annual Performance Evaluation Single Analyzer (PSD) | Quarterly | Percent difference of audit levels 3-10 < ± 15.1%, Audit levels 1&2 < ± 1.5 ppb difference or < ±15.1% |  |
| Federal Audit (NPAP) (SLAMS/NCORE) | 20% of sites audited in calendar year | Audit levels 1&2 < ± 1.5 ppb difference all other levels percent difference < ± 15.1% |  |
| Verification / Calibration | Upon receipt/adjustment/repair/ installation/moving Every 182 days and 2/calendar year if manual biweekly zero/span checks Every 365 days and 1/calendar year if continuous daily zero/span checks | Instrument residence time ≤ 2 minutes Dynamic parameter ≥ 2.75 ppm-min All points < ± 2.1% or < ± 1.5 ppb difference of best-fit straight line, whichever is greater and Slope 1 ± .05 |  |
| Gaseous Standards | All gas cylinders | NIST Traceable (e.g. EPA Protocol Gas) 50-100 ppm of NO in Nitrogen with < 1 ppm NO2 |  |
| Zero Air / Zero Air Check | Every 365 days and 1/calendar year | Concentration below LDL |  |
| Gas Dilution Systems | 1 / year or after failure of 1-point QC check or performance evaluation | Accuracy < ± 2.1% |  |
| Detection (FEM/FRMs) |  |  |  |
| Noise | Every 365 days and 1/calendar year | ≤ 0.005 ppm |  |
| Lower detectable level | Every 365 days and 1/calendar year | ≤ 0.01 ppm |  |

SYSTEMATIC CRITERIA – NO2

| Requirement (NO2) | Frequency | Acceptable Range | Review Comments |
| --- | --- | --- | --- |
| Standard Reporting Units | All data | ppb (final units in AQS) |  |
| Rounding convention for data reported to AQS | All data | 1 place after decimal with digits to right truncated |  |
| Completeness (SLAMS/NCORE/PSD) | Annual standard | ≥ 75% of hours in year |  |
| Completeness (SLAMS/NCORE) | 1-hour standard | Hour: ≥ 75% of hour Day: ≥ 75% of hours in a day Quarter: ≥ 75% sampling days in quarter Year: 4 quarters complete in each year 3 consecutive calendar years of complete data |  |
| Completeness (PSD) | 1-hour standard | Hour: ≥ 75% of hour Day: ≥ 75% of hours in a day Quarter: ≥ 80% of complete days Year: 4 consecutive quarters |  |
| Sample Residence Time Verification | Every 365 days and 1/calendar year | ≤ 20 seconds |  |
| Sample Probe, Inlet, Sampling train | All sites | Borosilicate glass,(e.g. Pyrex®) or Teflon® (FEP/TFE) |  |
| Pneumatic fittings | All tubing connections/valves | Teflon® (FEP/TFE), 316 stainless steel |  |
| Gas Regulators | All gas cylinders | 2-stage 316 stainless steel |  |
| Data Acquisition Systems | Digital or analog recording devices | Collection of continuous data(minimum of 1-minute values) |  |
| System Clock Verification | 1 / month | ≤ ± 1 minute |  |
| Siting | 1 / year | Meeting siting criteria or waiver documented, (PSD per approved QAPP) |  |
| Precision (using 1-point QC Checks) (SLAMS/NCORE) | Calculated Annually and as appropriate for design value estimates | 90% CL CV < 15.1% |  |
| Precision (using 1-point QC Checks) (PSD) | Calculated Quarterly | 90% CL CV < 15.1% |  |
| Bias (using 1-point QC Checks) (SLAMS/NCORE) | Calculated Annually and as appropriate for design value estimates | 95% CL < ± 15.1% |  |
| Bias (using 1-point QC Checks) (PSD) | Calculated Quarterly | 95% CL < ± 15.1% |  |
| Technical Systems Audits (SLAMS/NCORE) | 1 / 3 years | Confirmation of adherence to FRM/FEM, SOPs, and QAPP or documented waiver |  |
| Technical Systems Audits (PSD) | Annually (within 1 month of startup and annually thereafter) | Confirmation of adherence to FRM/FEM, SOPs, and QAPP or documented waiver |  |
| Annual PE Primary QA organization (PQAO) Evaluation | 1 / year | 95% of audit percent difference fall within the 1-point QC check 95% probability intervals at PQAO level of aggregation |  |