Critical Criteria PM10/PM2.5 Continuous

| Criteria (PM2.5/PM10) | 2) Frequency | 3) Acceptable Range | Comments |
| --- | --- | --- | --- |
| Data Reporting Period | Report every hour | A 24-hr period is calculated in AQS if 18 or more valid hours are reported for a day | 40 CFR Part 50 App N sec 3(c) |
| Average flow rate | Every 24 hour of ops; alternatively each hour can be checked | Average within 5% of 16.67 L/min at local conditions |  |
| Variability in flow rate | Every 24 hour of ops | CV ≤ 2% |  |
| One-point Flow Rate Verification | Every 30 days each separated by 14 days | < ± 4.1% of transfer standard & < ± 5.1% of flow rate design value |  |
| Design Flow Rate Adjustment | After multi-point calibration or verification | < ± 2.1% of design flow rate |  |
| External Leak check | Before each flow rate verification/calibration and before and after PM2.5 separator maintenance | < 1.5 L/min |  |
| PM10 inlet | At setup | Louvered PM10 size selective inlet as specified in CFR | 40 CFR App L, figs L-2 through L-19 |
| PM2.5 inlet | At setup | BGI VSCC or equivalent second stage separator approved for the method |  |

Operational Criteria PM10/PM2.5 Continuous

| Criteria (PM2.5/PM10) | 2) Frequency | 3) Acceptable Range | Comments |
| --- | --- | --- | --- |
| Temperature multi-point verification/calibration | On installation, then every 365 days and 1/calendar year | < ± 2.1° C |  |
| One-point temperature verification | Every 30 days | < ± 2.1° C |  |
| Pressure verification/calibration | On installation, then every 365 days and 1/calendar year | < ± 10.1 mm Hg |  |
| One-point pressure verification | Every 30 days | < ± 10.1 mm Hg |  |
| Flow rate multi-point verification/calibration | Following any electromechanical maintenance or transport, or every 365 days and 1/calendar year | < ± 2.1 % of transfer standard |  |
| 72-hour zero test | 1/yr. (1/6 mo. recommended) | Std dev of the data from a 72-hr zero test < 2.4 µg/m3 |  |

Precision

| Collocated samples | Every 12 days for 15% of sites by method designation | CV < 10.1% of samples ≥ 3 µg/m3 |  |
| --- | --- | --- | --- |

Accuracy

| Temperature audit | Every 180 days and at time of flow rate audit | < ± 2.1° C |  |
| --- | --- | --- | --- |
| Pressure audit | Every 180 days and at time of flow rate audit | < ± 10.1 mm Hg |  |
| Flow Rate audit | Twice a calendar year and 5-7 months apart | < ± 4.1 % of transfer standard & < ± 5.1 % of design flow rate |  |

Shelter Temperature

| Temperature range | Daily | 20 to 30° C (hourly avg.) or per manufacturer’s specifications if designated to wider temp. range (0 to +50° C for BAM 1020, shelter temp stable to within ± 2° C per hour) |  |
| --- | --- | --- | --- |
| Temperature control | Daily (hourly values) | < 2.1° C Std Dev over 24 hours |  |
| Temperature device check | Every 180 days and twice a calendar year | < ± 2.1° C |  |
| BAM routine inspection, cleaning, and maintenance |  |  |  |
| Inlet Head | Every 30 days | Inspect, clean if necessary |  |
| PM2.5 separator (VSCC) | Every 30 days | cleaned/changed |  |
| Down tube | Every 90 days | cleaned |  |
| Nozzle & vane | 1/mo. or more often as needed | Inspect and clean |  |
| Capstan shaft & roller | 1/mo. | Inspect and clean |  |
| Smart heater | 1/yr. | Inspect and maintain |  |
| Replace or clean pump muffler | 1/6 mo. | Inspect, clean or replace |  |
| Rebuild or replace pump | 1/yr. | Rebuild or replace |  |
| Clean or replace internal debris filter | 1/yr. | Clean, replace as needed |  |
| Membrane span foil check | 1/yr. | Avg. < ± 5 % of ABS value |  |
| Beta detector count rate | 1/yr. | Between 600,000 to 1,100,000 counts 4-min. test |  |
| Dark Count Value | 1/yr. | < 50 (recommended < 10)/4-min. test |  |
| Data comparison internal data logger to external data logger | Every 30 days 10 randomly selected values | Digital – exact match, analog ± 1 µg/m3 |  |

Systematic Criteria PM10/PM2.5 Continuous

| Criteria (PM2.5/PM10) | 2) Frequency | 3) Acceptable Range | Comments |
| --- | --- | --- | --- |
| Monitor | - | Meets requirements listed in the FEM designation |  |
| Siting | Every 365 days and 1/calendar year | Meets citing criteria or waiver documented |  |
| Data Completeness | Quarterly | ≥ 75% scheduled sampling days in each quarter |  |

Reporting Units

| PM2.5 | All concentrations | µg/m3 at ambient temperature & pressure |  |
| --- | --- | --- | --- |
| PM10L (local conditions) | All concentrations | µg/m3 at ambient temperature & pressure |  |
| PM10S (standard conditions) | All concentrations | µg/m3 corrected to standard temperature at 25° C & standard pressure at 760 mm Hg |  |
| Rounding convention for data reported to AQS | All concentrations | To one decimal place, with additional digits to the right truncated, or as reported by instrument |  |
| Annual 3-yr. average | All concentrations | Nearest 0.1 µg/m3 (≥ 0.05 round up) |  |
| 24-hour, 3-yr. average | All concentrations | Nearest 1 µg/m3 (≥ 0.5 round up) |  |
| Detection Limits |  |  |  |
| Lower detection limit | 24-hour avg. | < 1 µg/m3. |  |
|  | 1-hour avg. | < 4.8 µg/m3. |  |
| Upper detection limit | All hourly avgs. | 1000 µg/m3 |  |
| Verification/calibration standards recertification |  | All standards should have multi-point certifications against NIST-traceable standards |  |
| Flow rate transfer standard | Every 365 days and 1/calendar year | < ± 2.1% of NIST-traceable standard |  |
| Field thermometer | Every 365 days and 1/calendar year | ± 0.1° C resolution, ± 0.5° C accuracy |  |
| Field barometer | Every 365 days and 1/calendar year | ± 1 mm Hg resolution, ± 5 mm Hg accuracy |  |
| Clock/timer verification | Every 30 days | 1 minute/mo. |  |
| Data Precision |  |  |  |
| Single analyzer (collocated monitors) | Every 90 days | Coefficient of variation (CV) < 10.1% for values ≥ 3.0 µg/m3 |  |
| Primary Quality Assurance Org. | Annual and 3 yr. estimates | 90% CL of CV < 10.1% for values ≥ 3.0 µg/m3 |  |
| Data Bias |  |  |  |
| Performance Evaluation Program (PEP) | 5 audits for PQAOs with ≤ 5 sites 8 audits for PQAOs with > 5 sites | < ± 10.1% for value > 3µg/m3 |  |

Data Management Procedures and Documentation

| 1st Level Review | Comments and/or required actions |
| --- | --- |
| Daily review for limit checking, anomalies and irregularities |  |
| Review for data flags indicating out of limit conditions or instrument errors |  |
| PM coarse data comparison (PM10-PM2.5) for invalidation of negative values < - 5 µg/m3 |  |
| Logbook review: Station Logbook hardcopy and digital log entries |  |

| 1st Level Data Editing |  |
| --- | --- |
| Explanation of missing or invalidated data |  |
| Appropriate use of AQS null codes |  |
| Appropriate use of AQS data qualifiers |  |
| General Notes |  |

| Critical and operational QC forms and NIST Traceable Certifications | Archived in Air Tools. | Comments and/or required actions |
| --- | --- | --- |
| Most recent QA audit report and data sheets |  |  |
| 1-point QC forms for all leak, flow, pressure, temperature and time checks. |  |  |
| BAM multi-point verification/calibration data sheets |  |  |
| BAM 72-hour zero air background checks data sheets |  |  |
| Images of station logbook notes and/or digital record of electronic log entries |  |  |
| Images of current NIST traceable certifications for all reference devices and transfer standards. |  |  |
| Date & Performed by: |  |  |