Critical Criteria – PM2.5/PM10 Filter Based Local Conditions

Field Activities

| Criteria | Frequency | Acceptable Range | Review Comments |
| --- | --- | --- | --- |
| Sampler/Monitor | NA | Meets requirements listed in FRM/FEM/ARM designation | 40 CFR Part 58 App C 40 CFR Part 53 & FRM/FEM method list |
| Sampling Period |  |  |  |
| Sampling Period (including multiple power failures) | All Filters | 1380 – 1500 minutes, or if value <1380 and exceedance of NAAQS; midnight to midnight local std time | 40 CFR Part 50 App L (throughout) |
| Filter Holding Times |  |  |  |
| Sample Recovery | All Filters | ≤ 7 days 9 hours from sample end time |  |
| Pre-sampling | All filters | ≤ 30 days before sampling |  |
| Sampling Instrument |  |  |  |
| Average Flow Rate | Every 24 hours of operation | Average within 5% of 16.67 liters/minute |  |
| Variability in Flow Rate | Every 24 hours of operation | CV ≤ 2%; No flow rate excursions > ± 5% for > 5 minutes |  |
| One-point Flow Rate Verification | Every 30 days each separated by at least 14 days | ± 4.1% of transfer standard ± 5.1% of flow rate design value |  |
| Design Flow Rate Adjustment | After 1-pt or multi-point verification/calibration | < ± 2.1 % of design flow rate |  |
| External Leak Check | Before each flow rate verification/calibration and before and after PM2.5 separator maintenance | ≤ 8.5” Hg in 30 sec (2000H) ≤ 25 mm Hg in 60 sec (2000i) |  |
| Internal Leak Check | If failure of external leak check | ≤ 8.5” Hg in 30 sec (2000H) ≤ 140 mm Hg in 60 sec (2000i) |  |
| Filter Temp Sensor | Every 24 hrs of operation | No excursions > 5° C lasting longer than 30 minutes |  |

Laboratory Activities

| Criteria | Frequency | Acceptable Range | Review Comments |
| --- | --- | --- | --- |
| Post Sampling Weighing | All filters | Kept < 25˚ C from sample retrieval to conditioning; ≤ 10 days from sample end date if shipped at ambient temperature, or ≤ 30 days if shipped below avg. ambient (or < 4° C or below for avg. sampling temps < 4 °C) from sample end date |  |
| Filter Visual Defect Check (unexposed) | All filters | Correct type & size and for pinholes, particles or imperfections |  |

Filter Conditioning Environment

| Equilibration | All filters | 24 hours minimum |  |
| --- | --- | --- | --- |
| Temp. Range | All filters | 24-hr mean 20.0-23.0° C |  |
| Temp. Control | All filters | < 2.1° C Std Dev over 24 hours |  |
| Relative Humidity | All filters | 24-hr mean 30.0 – 40.0% RH or ≤ 5.0% sampling RH but ≥ 20.0% RH |  |
| Humidity Control | All filters | < 5.1% Std Dev over 24 hours |  |
| Pre/Post Sampling RH | All filters | Diff in 24-hr means < ± 5.1% RH |  |
| Balance | All filters | Located in filter conditioning environment |  |
| Balance auto-calibration | Prior to each weighing session | Manufacturer’s specs |  |

Operational Evaluations Table - PM2.5/PM10 Filter Based Local Conditions

Field Activities

Routine Verifications

| Criteria | Frequency | Acceptable Range | Review Comments |
| --- | --- | --- | --- |
| 1-point Temp. Verification | Every 30 days | < ± 2.1° C |  |
| Press. Verification | Every 30 days | < ± 10.1 mm Hg |  |

Annual Multi-point Verifications/Calibrations

| Temp. multi-point verifications/calibrations | Upon installation, then 1/yr. | < ± 2.1° C |  |
| --- | --- | --- | --- |
| Pressure verification/ calibration | Upon installation, and upon 1-pt verification failure | < ± 10.1 mm Hg |  |
| Flow rate multi-point verification/calibration | Upon installation, maintenance or transport, then 1/yr. | < ± 2.1% of transfer standard |  |
| Other Monitor Calibrations | Per manufacturer’s operation manual | Per manufacturer’s operation manual |  |

Precision

| Collocated Samples SLAMS | Every 12 days for 15% of sites by designation | CV < 10.1% of samples ≥ 3.0 µg/m3 |  |
| --- | --- | --- | --- |
| Collocated Samples PSD | Every 6 days for 15% of sites by designation | CV < 10.1% of samples ≥ 3.0 µg/m3 |  |

Accuracy

Note:  All equipment and transfer standards used for QA audits must be independent of the equipment used for routine QC activities such as 1-point and multi-point verifications and calibrations

| Temp. Audit | SLAMS every 180 days PSD Quarterly. | < ± 2.1 °C |  |
| --- | --- | --- | --- |
| Press. Audit | SLAMS every 180 days PSD Quarterly | < ± 10.1 mm Hg |  |
| Flow Rate Audit | SLAMS every 5-7 months PSD Quarterly | < ± 4.1% of audit standard < ± 5.1% of design flow rate |  |

Monitor Maintenance

| Very Sharp Cut Cyclone | Every 30 days | Cleaned/changed |  |
| --- | --- | --- | --- |
| Inlet/downtube cleaning | Every 90 days | Cleaned |  |
| Filter chamber cleaning | 1/mo. | Cleaned |  |
| Circulating fan/filter | 1/mo. | Cleaned/changed |  |
| Manufacturer’s Recommended Maintenance | Per manufacturer’s operations manual | Per manufacturer’s operations manual |  |

Laboratory Activities

Filter Checks

| Criteria | Frequency | Acceptable Range | Review Comments |
| --- | --- | --- | --- |
| Lot Blanks | 9 filters per lot | < ± 15.1 µg change between weighings |  |
| Exposure Lot Blanks | 3 filters per lot | < ± 15.1 µg change between weighings |  |
| Filter Integrity (exposed) | Each filter | No visual defects |  |

Lab QC Checks

| Field Filter Blanks | 10% or 1 per weighing session | < ± 30.1 µg change between weighings |  |
| --- | --- | --- | --- |
| Lab Filter Blanks | 10% or 1 per weighing session | < ± 15.1 µg change between weighings |  |
| Balance Check (working standards) | Beginning, 10th sample, end | < ± 3.1 µg |  |
| Duplicate Filter Weighing | 1 per weighing session | < ± 15.1 µg change between weighings |  |
| Microbalance Audit | 1/yr. | < ± 0.003 mg or manufacturer’s specs, whichever is tighter |  |

Verification/Calibrations

| Lab Temperature | Every 90 days | < ± 2.1 °C |  |
| --- | --- | --- | --- |
| Lab Humidity | Every 90 days | < ± 2.1 % |  |
| Microbalance Calibration | At Installation & 1/yr. | Manufacturer’s specifications |  |

Calibrations & Check Standards

| Working Mass Standards (compared to primary standards) Primary Standards | Every 90 days 1/yr. | < ± 2.1 µg 0.025 mg |  |
| --- | --- | --- | --- |

SYSTEMATIC CRITERIA – PM2.5/PM10 Filter Based Local Conditions

| Criteria | Frequency | Acceptable Range | Review Comments |
| --- | --- | --- | --- |
| Siting | 1/yr. | SLAMS Meets siting criteria or waiver documented PSD as per approved QAPP |  |
| Data Completeness SLAMS (3-year averaging period to calculate NAAQS compliance) | Annual Standard 24-hour Standard | ≥ 75% scheduled sampling days per quarter ≥ 75% scheduled sampling days per quarter |  |
| Data Completeness PSD (typically a 1-year monitoring period) | Annual Standard 24-hour Standard | ≥ 80% of scheduled sampling days per quarter ≥ 80% of scheduled sampling days per quarter |  |
| Reporting Units | All filters | µg/m3 at ambient temp/press (PM2.5) |  |
| Rounding convention for DV calculation and data reported to AQS | All filters | To one decimal, with additional digits to the right being truncated |  |
| Annual 3-yr average | All concentrations | Nearest 0.1 µg/m3 (≥0.05 round up) |  |
| 24-hour, 3-yr average | All concentrations | Nearest 1 µg/m3 (≥0.5 round up) |  |
| Detection Limit |  |  |  |
| Lower detection limit | All filters | ≤ 2 µg/m3 |  |
| Upper concentration limit | All filters | ≥ 200 µg/m3 |  |
| Precision |  |  |  |
| Single analyzer (collocated monitors) | Every 90 days | Coefficient of variation (CV) < 10.1% for values ≥ 3.0 µg/m3 |  |
| Primary Quality Assurance Organization (PQAO) | Annual and 3 year estimates | 90% confidence level (CL) of CV < 10.1% for values ≥ 3.0 µg/m3 |  |
| Bias |  |  |  |
| Performance Evaluation Program (PEP) SLAMS | 5 audits for PQAOs with ≤ 5 sites 8 audits for PQAOs with > 5 sites | < ± 10.1% for values ≥ 3.0 µg/m3 |  |
| Performance Evaluation Program (PEP) PSD | 1/yr. | < ± 10.1% for values ≥ 3.0 µg/m3 |  |
| Technical Systems Audit SLAMS | 1/3 yr. | Review of entire field, lab, and data reporting process for comparison to QC/QA requirements |  |
| Technical Systems Audit PSD | 1/project, if extended 1/yr. | Review of entire field, lab, and data reporting process for comparison to QC/QA requirements |  |

Field Activities

Verification/Calibration Standards Recertification – All standards should have multi-point certifications against NIST traceable standards

| Criteria | Frequency | Acceptable Range | Review Comments |
| --- | --- | --- | --- |
| Flow Rate Transfer Standard | Every 365 days | < ± 2.1% of NIST Traceable Standard |  |
| Field thermometer | Every 365 days | ± 0.1° C resolution ± 0.5° C accuracy |  |
| Field barometer | Every 365 days | ± 1 mm Hg resolution ± 5 mm Hg accuracy |  |
| Clock/timer verification | Every 30 days | ± 1 min NIST AST |  |

Laboratory Activities

| Criteria | Frequency | Acceptable Range | Review Comments |
| --- | --- | --- | --- |
| Microbalance Readability | At purchase | 1µg |  |
| Microbalance Repeatability | At purchase | 1µg |  |
| Primary mass/Working mass Verification/Calibration Standards Recertification | 1/yr | 0.025 mg tolerance |  |